| Supplier: | HTI | Type Nr.: |  |
| --- | --- | --- | --- |
| Product: | Smart Suction cup | Art. no.: |  |
| Project Number | TBC | Issued by PM: | Mandar / David |
| EE/ME | Danny/Zenger | QE/ Compliance | Tim/Ice |
| Any change on the product, part or process as well as change of sub supplier has to be approved and confirmed by Therabody in written form. |  |  |  |

|  |  | Product Specification PM |  |  |
| --- | --- | --- | --- | --- |
| Pictures / Design: |  |  |  |  |
| Replacement for: |  |  |  |  |
| 1.1 | Specification / Features |  |  |  |
| Easy portable |  |  | Yes, attached & de-attached at least 4500 times in total |  |
| Vibration |  |  | Yes, 1 motor |  |
| Battery |  |  | Yes |  |
| Heating element: |  |  | Yes |  |
| Control Panel |  |  | Yes, integrated into the device |  |
| Bluetooth ready |  |  | Yes, remote control range >15m |  |
| Vibration area |  |  | TBC |  |
| Washable/water proof |  |  | Yes, cup is washable/device is cleaned with wet cloth, IP22 for medical device |  |
| Air Pump |  |  | Yes |  |
| Hook and rope |  |  | TBD |  |
| Cup |  |  | Quantity 3 cups TBC |  |
| 1.2 | Optical / Design / Material Specification |  |  |  |
| Decoration sheets |  |  | No, but cosmetically appealing |  |
| Charging time |  |  | <70mins 1C charging |  |
| Discharging time |  |  | 120mins with full load |  |
| Display functions |  |  | No display, only LED indicators |  |
| LED indicator |  |  | 3 modes 9 LEDs, timer 3 LEDs, battery energy 1 RGB LED Refer to software spec |  |
| Control Panel material |  |  | TBC (HTI) |  |
| Packaging |  |  | acc. to Therabody design TBC: 1 device with 3 size cups, 3 devices with 3 size cups, 3 devices with 9 cups(3 size) |  |
| IM |  |  | Instruction manual acc. to Therabody design |  |

|  | Technical requirements |  |  |  |
| --- | --- | --- | --- | --- |
| 2.1 | General Technical requirements |  |  |  |
| Rating: |  | 3.7V, max. 1.4A | Dimensions (L x W x H) [mm]: | D67mmx95mm, |
| Weight: |  | Approx. 220g | Noise: | TBC, should be smaller than 50dB |
| Temperature rise time: |  | Target 30secs to 43C | Working temperature: | -10C to 40C |
| Heating temperature setting: |  | TBC | Working duration time: | 5, 10 or 15mins per cycle |
| Working time with full loading: |  | 120mins | Battery capacity: | 2000mAh, Support 1C charge and 2C discharge. |
| Charge time: |  | <70mins | Battery size: | Size:18550 (HTI) |
| Charging port: |  | USB-C | USB cable: | USB-A to USB-C and 1m |
| MCU voltage ripple: |  | <60mV | EMC margin: | >5dB |
| Button debounce time: |  | 50mS | Sleep current (Standby): | <40uA |
| Temperature protection: |  | Yes, NTC + software Stop heating if NTC detect normally. (HTI) | Air Pump max pressure: | >60KPa |
| Vibration intensity: |  | TBC RMP:900, 1800, 3000, Hz: 15, 30, 50 (HTI) | Vibration area: | 35mm, 45mm, 55mm |
| Suction intensity: |  | -30Kpa, -45Kpa, -60Kpa | FDA Grade: | US/EU/AU/UK Class I, JP/CA Class II? |
| Pressure relief valve: |  | Yes, ?Kpa (HTI) | Treatment mode: | Pressure+vibration+heat Pressure mode: Breath suction or steady suction US team |
| Drop test: |  | 1.2m height, cement ground | Storage temperature: | According to TTR temperature test requirement |

Battery level RGB indicator LEDs definitions and setting.

Charging mode battery LED indicator：The RGB LED indicator definitions.

|  |  | RGB LED indicator | Battery capacity |
| --- | --- | --- | --- |
| Battery charging | Red LED breathing |  | 0~25% |
| 2 | Blue LED breathing |  | 26%~75% |
|  | Green LED breathing |  | 76%~99% |
|  | Green LED solid on |  | 100% |

Remark: Breathing patten is 2S ON,0.5S OFF. (Fading)

Non-charging mode battery LED indicator. The RGB LED indicator definitions. (Short press the suction button to display the battery capacity. The device auto turns off after 3mins.

|  | RGB LED indicator | Battery capacity |
| --- | --- | --- |
| Device | Red LED flashing | <15% |
|  | Red LED solid on | 15~25% |
|  | Blue LED solid on | 26%~75% |
|  | Green LED solid on | 76%~100% |

Remark: When reach 0%，red LED flash by 5 times then device turns off.

Bluetooth indicator LEDs definitions and settings.

- Press any button, it will wake up the battery level LED for 5secs and the blue tooth keep been searching for 3mins for user connecting to the APP.

- After connecting to the APP, all LEDs including battery level LED and the speed or mode LEDs will flash with 1Hz.

- After 3mins no activation, the device sleep and all LEDs turn off.

Firmware and test APP requirements:

- The blue tooth name can be changed by user.

- Follow Therabody blue tooth communication protocols format.

- Has manufacturing/MCU code in firmware for APP team tracking.

- The firmware can be upgraded and downgraded.

- Test app can check the CRC code for the firmware. The firmware only can be updated when the CRC code check correct.

### Set product manufacturing information：

a	Host sends command to Slave

| Start | Checksum | Command | Data Length |
| --- | --- | --- | --- |
| 0xAA | 0x20 | 0x20 | 0x00 |

b	Slave responses to Host

| Start | Checksum | Command | Data Length | D0 | D1 | D2 | D3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0xBB | 0x xx | 0x20 | 0x10 | Delivery time HH | Delivery time H | Delivery time L | Delivery time LL |
| D4 |  | D5 |  | D6 | D7 | D8 | D9 |
| Manufacturer code |  | Product code |  | MCU code | reserved | Firmware Version H | Firmware Version L |
| D10 |  | D11 |  | D12 | D13 | D14 | D15 |
| Communication protocol version H |  | Communication protocol version L |  | APP change product parameters setting Bit0: bluetooth name | reserved | Product type number H* | Product type number L* |

D4 list(Manufacturer code): 0x03(HTI)

D5 list(Product code):

| Item | Product type name | Bluetooth name | Product code(hex) | Remark |
| --- | --- | --- | --- | --- |
| 9 | Smart suction cup | Smart suction cup | 160(0xA0) | Support D0~D15 |

D6 list(Product code):

| Item | Products | MCU | MCU code(hex) | Remark |
| --- | --- | --- | --- | --- |
| 1 | 1st MCU | Compatible | 1(0X01) | Support D0~D15 |

|  | Workmanship |  |  |
| --- | --- | --- | --- |
| 3.1 | Handling / function ability |  |  |
| Buttons / knob must work smoothly. |  |  |  |
| Easy to use and self-explanatory operation. |  |  |  |
| Easy cleaning. |  |  |  |
| No sharp edges allowed. |  |  |  |
| All plastic – and lacquered parts must be UV resistant |  |  |  |
| The item must be BPA free (limit 0 mg/kg) |  |  |  |
| A permanently bad smell during normal use is not acceptable |  |  |  |
| No big gaps and height differences between housing parts allowed. Gaps and height differences must be even |  |  |  |
| Rating label must be reliable to N-Hexane (15 sec rubbing test) |  |  |  |
| 3.2 |  | Material / colour / coating |  |
| Mat |  |  | PC+ABS, color TBC |
| Control panel |  |  | Material TBC, color TBC |
| Adapter |  |  | Not included. |
| Decoration film/layer |  |  | No |
| Etc. |  |  |  |

|  | Product safety & documents |
| --- | --- |
| The seller has to ensure, that all documents are available, that Therabody can start the certification at an independent institute. |  |
| Documents from independent certification body: |  |
| Certificate and report with country deviation from an independent institute. Tests acc. to IEC and EN safety standard) |  |
| EMC certificate and report acc. to IEC/EN standard |  |
| RoHS |  |
| REACh |  |
| Bill of Material (BOM) |  |
| Explosion Drawing |  |
| Electric circuit diagram, layout drawing, firmware bin file (including CRC code or checksum) |  |
| Rust test acc. to Therabody Humidity Test Report (HTR) |  |
| UV stability test for plastic and painted parts |  |
| Supplier LTT (life time test) report |  |
| Packaging test report from an independent institute |  |
| etc. |  |

|  | Long Life Test Requirements |  |
| --- | --- | --- |
| Requirements according to: |  | Therabody |
| Life time: |  | 2 years / 1200 hours |
| Cycles / hours |  | N/A |
| Cycle description: |  | One cycle: 15 mins ON, 15 mins Off |
| Load: |  | Full load |
| Test condition: |  | 20-25 degree, ambient temperature |
| Test description: |  | N/A |
| Approx. testing time: |  | 1200 hours running time |
| Requirement: |  | All tested samples must pass the life time test. During testing no visual or functional changes of the samples are allowed. The requirement of the standard must be fulfilled anyway. |

| 6. | Revision History |  |
| --- | --- | --- |
| Status | Date | Change |
| 1 | 08.09.2021 | First Draft |
| 2 |  |  |
| 3 |  |  |